## Slide 1
Accelerator Mass Spectrometry (AMS)
496. Wilhelm und Else Heraeus-Seminar
February 6th to 10th, 2012 – Physikzentrum Bad Honnef – Germany
Astrophysics with modern small-scale accelerators

## Slide 2: Outline
Method – Accelerator Mass Spectrometry (measure rare isotopes)
AMS at Munich (Garching)
New Facility DREAMS (DREsden AMS)
Notes: Rare isotopes: radioactive ones – other elements, isotpes are everywhere

## Slide 3: Historical
1939 Alvarez and Cornog discovered accidentally 3He (cyclotron).
From 1939 until 1977, for around 40 years, dark and silent!
1977 Muller proposed to measure 14C or 10Be by means of a cyclotron.
Independently, Gove, Litherland, and Purser suggested the use
of a Tandem accelerator for 14C measurements.
2012 More than 100 AMS laboratories worldwide.
From small table top accelerators until large facilities with more than 10MV terminal voltage.
Isotopes of interest from 3H until the heaviest nuclides like 244Pu.
The majority of AMS laboratories is dedicated to 14C and 10Be.

## Slide 4: Accelerator Mass Spectrometry
Counting individual atoms (like conventional MS)
But
accelerator (mostly a Tandem)
molecules destroyed by stripping
mostly for radioisotopes
extremely sensitive (isotopic ratio ~ 10-16)
Notes: This sensitivity is good but extremely sensitive means that all stable isotopes are in your sample. Therefore not all isotopes are suited for this method
We need isotopes which are produced and which decay – radionuclides (next slide).
4

## Slide 5: Radionuclides detected by AMS
Fig. 4 from D.Fink, NIM B 268, 2010 modified/updated from W. Kutschera 1981
Notes: Here is an overview of radionuclides measured (successfully) by AMS – mass – half life – very long half life – primordial elements like stable elements still abundant shorter half life suited for AMS – growing number of isotopes measured with the years – short half lives – you can measure the activity. So what are the challenges at AMS?
Like 44Ti where decay counting is also possible (T1/2: ~60 a)
5

## Slide 6: Accelerator Mass Spectrometry
challenge: isobar separation
luck (e.g. for 14C- no stable 14N- ion)
X-ray coincidence
chemistry
specific energy loss
deflection in a gas-filled magnet
Example (63Ni in Cu):
m(63Ni) = 62.929674 (T1/2 = 100 a)
m(63Cu)= 62.929602 (stable)
Dm: < 1.2 10-6
At Munich/Garching DREAMS later
Notes: Again challenge of AMS – isobar separation!
Different methods – most famous –wide applications for AMS C-14 luck – understanding the whole process many disciplines involved to make the ultrasensitive measurements
As an example – worst case looking for an radionuclide in a matrix of the isobar!

## Slide 7: Garching Research Center about 14km North of Munich

## Slide 8: Tandem Accelerator Laboratory Garching Part of the MLL (Maier Leibnitz Laboratory) Ludwig Maximilians Universität and Technische Universität München
Length: about 76 m

## Slide 10: AMS Setup in Garching
Wien filter 2
1
Notes: Start negative ion source, first mass spec, molecular background, high voltage tandem accelerator, stripping foil destruction of molecules, Wien filter (1 and 2) to separate the right velocity,2nd mass spec to separate a charge state, beam line to filter out background originating from scatter in the beam – guided to the detector system GAMS – I will explain the detector briefely.
10

## Slide 11
Neg. Cs – Sputter Ion Source

## Slide 12: AMS Setup in Garching
Wien filter 2
1
Notes: Start negative ion source, first mass spec, molecular background, high voltage tandem accelerator, stripping foil destruction of molecules, Wien filter (1 and 2) to separate the right velocity,2nd mass spec to separate a charge state, beam line to filter out background originating from scatter in the beam – guided to the detector system GAMS – I will explain the detector briefely.
12

## Slide 13
Low energy side of the Tandem

## Slide 14
e.g.: 63Cu12+ + 63Ni12+ isobars not yet suppressed
High energy side of the Tandem

## Slide 15: AMS Setup in Garching
Wien filter 2
1
Notes: Start negative ion source, first mass spec, molecular background, high voltage tandem accelerator, stripping foil destruction of molecules, Wien filter (1 and 2) to separate the right velocity,2nd mass spec to separate a charge state, beam line to filter out background originating from scatter in the beam – guided to the detector system GAMS – I will explain the detector briefely.
15

## Slide 16
Wien Filters
WF off
WF on
60Ni
59Co
59Ni
Mass resolution: distance between peaks d ≥ FWHM
d
FWHM
Notes: This slide shows the effect of the velocity filters when measuring highly sensitive 59Ni/60Ni ratio. The uninteresting isotope (60Ni) is efficiently supressed with the correct setting of the Wien filter. It is important to have alternating electric and magnetic separation. Therefore two new Wien filters were put into operation.

## Slide 17: AMS Setup in Garching
Wien filter 2
1
Notes: Start negative ion source, first mass spec, molecular background, high voltage tandem accelerator, stripping foil destruction of molecules, Wien filter (1 and 2) to separate the right velocity,2nd mass spec to separate a charge state, beam line to filter out background originating from scatter in the beam – guided to the detector system GAMS – I will explain the detector briefely.
17

## Slide 18: The Gas-filled Analyzing Magnet System
Notes: 18

## Slide 19: Isobar Separation in a Gas-filled Magnet
magnetic rigidity
in a gas-filled magnet
Notes: Separation in a GFM – magnetic rigidity can be seen in this formula – if we have a gas the incoming beam interacts with the gas atoms and the charge changing reactions lead to mean charge state – then deflection with a magnet -> separation in position -> place the detector at the right position – block the isobar!
19

## Slide 20
e1l
e1r
e2l
e2r
e3
e4
e5
Frisch grid: Etot
y-angle: dt
9 „independent“ signals
position: p
x-angle: dp
Ionization chamber
dE
Notes: 20

## Slide 21: spectra for calibration sample: 60Fe/Fe = 5 ·10-12
position
position
position
position
position
position
Etot
dE1
dE2
dE3
dE4
dE5
60Fe
60Fe
60Fe
60Fe
60Fe
60Fe
60Ni
Ni/Fe ≈ 10-5 => 60Ni/60Fe ≈ 106

## Slide 22: spectra for crust sample: 60Fe/Fe ~ 2·10-15
position
position
position
position
position
position
Etot
dE1
dE2
dE3
dE4
dE5
60Fe
60Fe
60Fe
60Fe
60Fe
60Fe
with windows
60Ni

## Slide 23: spectra for crust sample: 60Fe/Fe = 1.4 ·10-15
position
position
position
position
position
position
Etot
dE1
dE2
dE3
dE4
dE5
60Fe
60Fe
60Fe
60Fe
60Fe
60Fe
with windows

## Slide 24: some numbers
1 mA 56FeO- => 6·1012/sec
transmission 5% => 3·1011/sec
for 60Fe/56Fe = 10-15 => 3·10-4/sec = 1event/hour
10 mg 56Fe => 1020 atoms
1 mA 56FeO- => 6·1012 atoms/sec
neg. ion yield 10-3 => 6·1015 atoms/sec used
sample lasts for => 1.5·104 sec = 4 hours

## Slide 25
Application: Manganese Crusts
Notes: Was suchen wir in Mangankrusten? Welches Signal? Nicht historische Vergangenheit – ist eher ungewöhnlich für Messungen mit Radioisotopen mit langer Halbwertszeit. – ich versuche das zu illustrieren:

## Slide 26: Reservoir: Manganese Crust

## Slide 27
10Be
10Be, T1/2= 1.387 Myr
Korschinek et al. NIMB 2010
Chmeleff et al. NIMB 2010

## Slide 28: Sample Preparation
Layers of 1-2 mm
Chemical Preparation
1: Whole material
2: Leaching of Fe
Mixing with Ag powder
In Sample holder:

## Slide 29
(Cs Sputter)
Negative
Ion Source
Sample
90° Injector Magnet
(1st Mass Separation)
18° Electrostatic
Deflection
Carbon
Stripper
Foil
S
v>vi
v<vi
v<vi
+
−
N
N
−
+
v>vi
m<mi
m>mi
Wien
Filter
Wien
Filter
90° Analyzing Magnet
(2nd Mass Separation)
Switching Magnet
Gas-filled
Analyzing
Magnet
System
135° Magnet
B= 0.6-1.2 T
p= 3.5-7 mbar N2
Ionization Chambers
S
Faraday
Cup


m>mi
m<mi
Accelerator Mass Spectrometry
at the
14 MV Munich MP Tandem
Negative
Ions
Positive
Ions
3.19m
TOF
Start
TOF
Stop
−
+
Wien
Filter
Time-of-flight setup
(3.19m flight path)


Thanks to Iris Dillmann

## Slide 30: Measurements with GAMS: Deep See Ferro-Manganese Crust
Knie et al. PRL 93 (2004)

## Slide 31
The pacific was enriched in 60Fe 2.2 Myr ago!
Leaching of the Ferro-Manganese Crust
Fitoussi et al. PRL 93 (2008)

## Slide 32: Signal in more reservoirs: Moon
collaboration with G. Herzog et al. (Rutgers University) and NASA
First measurement!
D.L. Cook et al., 40th Lunar and Planetary Science Conference (2009), 1129.
Notes: From lunar compendium meyer et al. 2007 A16drill.pdf
Figure 1: Photo of ALSEP site at Apollo 16 after deep drill core was taken. The drill moter
is in view. Approximate location of drill hole in red. Stone Mountain about 5 km distant.
AS16-113-18367.

## Slide 33
First measurements: 60Fe from Lunar Samples
Cook et al, LPSC XXXIX (2009)
Cook et al, LPSC XXXIX (2009)
Potential origins:
  Galactic Cosmic Rays
  Solar Cosmic Rays
  Deposition of SN debris
P60Fe-GCR = (0.88±0.44) dpm/kg Ni
P60Fe-SCR = 0.07– 0.45 dpm/kg Ni on the surface
Knie et al, Planetary Science Letters 34, 729 (1999)
Cook et al, LPSC XXXIX (2009)
Notes: This graph show the results of 60Fe measurements performed in the year 2005 in the Maier Leibnitz Laboratorium also by our group. The graph on the left show the 60Fe to stable iron. Signal of 60Fe above blank level were observed in all five samples from the core of Apollo 16 and in the uppermost sample from core 12025/8, but not in any of the deeper samples from this core, where no single 60Fe event was detected and only upper limits could be determined.
There are 3 potential origins for 60Fe on the Moon: Production by galactic cosmic rays, production by solar cosmic rays or deposition of SN debris. We can estimate the contribution of GCR by comparing our results with the activity measured in iron meteorites. GCR produce a mean rate of 0.88 dpm/kg Ni (solid line on the right graph, dashed line indicate 1 SD deviation). All five sample from Apollo 16 have activities that are comparable to the range of GCR values in meteorites. Almost all of the other samples had no detectable 60Fe, but the large uncertainties place these results within the range of GCR values.
Calculation have been performed in order to estimate the production rate by SCR. This is greatest at the surface and quickly decreases with depth. No contribution is expected at the depth of the samples from the Apollo 16 core.

## Slide 34: AMS at GAMS
Isotope | application | ion in ion out | isobar | isobar Supp. | sensitivity
26Al | astroph. x-section manganese crusts | AlO- Al7+ | 26Mg | 109 | few 10-15
36Cl | groundwater, meteorites | Cl- Cl10+ | 36S | 109 – 1010 | < 10-15
41Ca | Hiroshima teeth, meteorites | CaH3-,CaF3- Ca9+ | 41K | 109 – 1010 | < 10-15 10-14 CaF
44Ti | astroph. x-section | TiO- Ti9+ | 44Ca | 109 – 1010 | few 10-15
53Mn | crusts, in-situ, meteorites, sediments, moon rocks | 53MnO- 53Mn11+ | 53Cr | 109 | 10-14 – 10-12
59Ni | nuclear waste, meteorites, astroph. x-section (NIMB2007) | 59Ni- 59Ni12+ | 59Co | 109 | few 10-14
60Fe | crusts, meteorites, moon rock, nuclear waste | 60FeO- 60Fe11+ | 60Ni | >1011 | few 10-16
63Ni | Hiroshima copper, neutron dosimetry, astroph. x-section | 63Ni- 63Ni12+ | 63Cu | 108 – 109 | few 10-14
79Se, 93Mo,93Zr, 107Pd | astroph. x-section (Proc. NiC IX 2006), half-life, nuclear waste | various | 79Br, 93Nb, 107Ag | 106 | few 10-12 ? few 10-10 ? ?

## Slide 35: AMS Setup in Garching
Andres Arazi
Pankaj Baghel
Shawn Bishop
Iris Dillmann
Thomas Faestermann
Leticia Fimiani
Jose M. Gomez Guzman
Karin Hain
Klaus Knie
Gunther Korschinek
Johannes Lachner
Peter Ludwig
Moumita Maiti
Astrid Meier
Hisham Nassar
Mikhail Poutivtsev
Werner Rühm
Georg Rugel
Anton Wallner
Wien filter 2
1

## Slide 36: Accelerator Mass Spectrometry
challenge: isobar separation
luck (e.g. for 14C- no stable 14N- ion); X-ray coincidence; deflection in a gas-filled magnet
chemistry
specific energy loss
absorber foil and ESA (electrostatic analyzer)
DREAMS
Example (10Be and 10B):
m(10Be)= 10.0135 AMU (T1/2 = 1.387 Myr)
m(10B) = 10.0129 AMU (stable)
Dm: 6 × 10-5
Notes: Again challenge of AMS – isobar separation!
Different methods – most famous –wide applications for AMS C-14 luck – understanding the whole process many disciplines involved to make the ultrasensitive measurements
As an example – worst case looking for an radionuclide in a matrix of the isobar!

## Slide 37
AMS ion sources
AMS detection system
Dresden AMS facility - DREAMS

## Slide 38
Ion Beam Center at Helmholtz-Zentrum Dresden-Rossendorf
Chambers
Magnets
Accelerators
Ion sources
Beam lines
10 m
New 6 MV accelerator

## Slide 39
6 MV Tandetron for AMS, IBA and HEII
6 MV Tandetron
AMS section
Ion implantation
section
AMS injector
IBA & II injector
AMS sputter sources
90°- AMS analyzing magnet
Faraday cups for 9Be, 12C, 13C, 27Al, 40Ca, 35Cl, 37Cl, 127I, …
Absorber foil for isobar suppression
35° ESA
30° vertical magnet
Gas ionization detector
Automatic
wafer handling system
for HE ion implantation
IBA analyzing magnet
10 m
Terminal voltage: 0.3 - 6.0 MV
Maximum TV: 6.6 MV
Terminal voltage stability: ~ 10-5
Installed in 2009-2010

## Slide 40
AMS ion source:
Cs sputter ion source SO-110 (HVEE)
Insulator
Insulator
Extractor
Negative Ions
Ionizer
Ionizer shroud
Cs - Vapor
Cathode
Magnetic
manipulator
(HV insulated)
20 - 30 kV
Extraction voltage
3 - 10 kV
Sputter voltage
Cooled base
AMS target holder
1 cm
Wheel for 200 samples (4 x 50)
20 cm

## Slide 41
54° ESA
AMS-Source
Bouncer magnet
X-Slit
AMS ion injector
Magnet chamber
on high potential: up to 3 kV
Switching frequency: ~100 Hz
stable nuclide: 100 µs
Radionuclide: 10 ms
Waiting: 400 µs
Better precision by an
sequentially isotope injection by a bouncer magnet

## Slide 42
6 MV Tandetron
AMS section
Ion implantation
section
AMS injector
IBA & II injector
AMS sputter sources
90°- AMS analyzing magnet
Faraday cups for 9Be, 12C, 13C, 27Al, 40Ca, 35Cl, 37Cl, 127I, …
Absorber foil for isobar suppression
35° ESA
30° vertical magnet
Gas ionization detector
Automatic
wafer handling system
for HE ion implantation
IBA analyzing magnet
10 m
Terminal voltage: 0.3 - 6.0 MV
Maximum TV: 6.6 MV
Terminal voltage stability: ~ 10-5
6 MV Tandetron for AMS, IBA and HEII

## Slide 43
Quadrupole
triplet lens
35°-elektrostatic analyzer (ESA)
Aperture 4
90°-AMS magnet
Faraday cup chamber
Moveable FC
Fixed FC (for 12C)
Quadrupole lens
Quadrupole lens
30° vertical magnet
Detector (Gas ionization chamber)
Faraday cup
Absorber foils
AMS facility: high energy side

## Slide 44
90°-AMS magnet
Radioisotope path
to absorber foil and 35°-ESA
2 moveable Faraday cups
High precision moveable FC for stable isotopes and energy stabilization
Slit system for HV stabilization in the moveable FCs
AMS facility: measurement of stable isotopes
fixed Faraday cup
Faraday cup chamber
Fixed FC
Moveable FC

## Slide 45
Foil thickness: 1 µm, 500 nm, 200 nm
Isobar suppression
Main interest at DREAMS: 10Be, 26Al, 36Cl, 41Ca and 129I
10Be - 10B
absorber foil and 35° ESA for isobar suppression
36Cl - 36S
Si3N4- absorber foil holder
Si3N4- absorber foil set
Faraday cups
35°- ESA
Q-lens
Beam

## Slide 46
Vertical 30°-magnet and gas ionization detector
Gas: Isobutene (8 - 40 mbar)
4 anodes: ΔE1, ΔE2, ΔE3, Erest
Window: 75 nm thick silicon nitride (8x8 mm2)
AMS detector

## Slide 47
10 cm
HVEE
Detector without housing
36Cl
36S
Isobutene, 17.5 mbar
Gas ionization detector with 4 anodes
A B C D Anodes
Ions
Isobutane C4H10
AMS detector
Ions
Gas ionization

## Slide 48
Test with 10Be samples
current dependence: NIST 4325 standards

## Slide 49
Replica measurements of 10Be isolated from material originating from the LL chondrite Dhurmsala
Notes: Fig. 3: Replica measurements of 10Be isolated from material originating from the LL
chondrite Dhurmsala (Fragment 2/3A, ~70 g without ~1 g of metal, homogenised to size fraction of <125 μm [NEU93]. AMS targets have been prepared at four chemistry labs (U of Cologne, U Hanover, Rutgers U, PRIMELab) by eight scientists (M. Altmaier, W. Klas, M. Knauer, B. Meltzow, S. Merchel, J. Mokos, U. Neupert, C. Schnabel) [MER98 and references therein]. AMS measurements have been performed from 1994-1998 at ETH Zurich (values adapted to new half-life and standard values [KUB10]) and PRIMElab. The new DREAMS 10Be value is from BeO prepared in 2011 from ~150 mg of the same Dhurmsala fraction.

## Slide 50
total transmission (%)
currents (µA)
terminal voltage (MV)
stripped to
(with absorber foil)
extraction as
total with absorber foil
only absorber foil
suppression factor
background
precision (%)
10Be
26Al
36Cl
41Ca
129I
Be2+ Al3+ Cl5+ Ca4+ I5+
Be4+ Cl11+
4.5 2.7 6.0 5.8 5.0
2 - 7 0.2 - 0.5 20 - 30 0.15 - 0.2 3 - 5
BeO- Al- Cl- CaF3- I-
58 35 26 16 9
36
12
21
3
4.5·10-16 8·10-16 3 ·10-15 8 ·10-15 2 ·10-13
0.7 2.3 2.7 2.5 1.2
(at 10-12) (at 10-12) (at 10-12) (at 10-11) (at 10-11)
2·106
7·103
*no AgBr backing! Ni cathodes with Ni pins
Current performance data DREAMS (2010/2011)

## Slide 51
Thanks to the AMS group at HZDR (DREAMS)
Sh. Akhmadaliev,
S. Pavetich, S. Merchel

## Slide 52
New Isotopes ?
First: Look into the
Ion Cook Book
http://tvdg10.phy.bnl.gov/ COOKBOOK/
Scanned version available under:
Find someone for your DREAMS to become true !

## Slide 53
Thank you for your attention !